MEMO

7600E 4-20ma / 0-10 V Analog Setup

- Press and hold the ZERO button to enter programming mode

- Display will show “P     XXX” then “C    XXX” then “ENT AC”

- Key in four 0’s

- Press ENT

- Display will read “SEL CFG”

- Use the TARE RECALL button to select “SEL.OP 1”

- Press GROSS NET button to enter Option 1 (analog setup)

- Display will show “1.1    Gr”

- Use the TARE RECALL button to select if the analog option is to follow the gross wt (Gr), the net weight (nEt), or the displayed weight (dSP)

- Press GROSS NET

- Display will read “1.5   Zr” and 0.00. This is the weight at which the analog output is to be 0 volts or 4 ma. Key in the desired number and press ENT

- Press GROSS NET button

- Display will read “1.6   FS” and the full capacity of the scale. This is the weight value at which the analog output will read 10 volts or 20 ma. Key in the desired number and press ENT

- Press GROSS NET button

- Display will read “1.7  ZR.A”. Fine adjustments may be made to the analog output zero point at this step. Hook a volt meter or amp meter to TB 20 and use the TARE RECALL button to increase the output and the START STOP button to decrease the output

- Press GROSS NET button

- Display will read “1.6  SP.A” Fine adjustments may now be made to the  analog output full scale point at this step. Hook a volt meter or amp meter to TB 20 and use the TARE RECALL button to increase the output and the START STOP button to decrease the output

- Press ENT button

- Display will show “SEL.OP 1”

- Press ENT

- Display will show “SAVE ?” and “no”

- Use the UNITS button to select “YES”

- Press ENT

- Display will show “ENT AC”

- Press ENT and the 7600E will come out of the programming mode